## Defibrillator
Lot No., Item No. | [6 empty] | Bidder's reference:
CA Reference |  |  |  |  | GEN20/109741D/0001 | Revision No.
CA Specification | [5 empty] | Revision Date
Prepared by |  |  |  |  | HK | Bidder's reference:
DETAILED SPECIFICATION | [6 empty] | STATEMENT of COMPLAIANCE (*)
Instructions to Bidder: Bidders are to offer a standard production model most closely matching the specification below and provide details of the offer. * Bidders must enter "Comply" or "Not comply" against each paragraph of this specification, comment as necessary, and sign and stamp each page. Where information is presented in the specification, the comment should be “Noted and Understood”. Failure to complete this statement of compliance may result in the offer being rejected. In the event of electronic bidding by E-mail or by means of scanning of the Specification for completion by word processing, the specification shall not be altered in any way. Any deviation to specification should be stated and if necessary, fully explained as a comment in the compliance column without making any changes to the specification. Any changes to the specification may invalidate the Bid. Each compliance sheet shall be duly authorised, signed, dated and stamped using the bidders company stamp. | [6 empty] | UNITS ARE TO BE STATED USING THE SI SYSTEM. For example length in metres (m). Multiples and sub-multiples to be used are: 106 mega (M); 103 kilo (k); 10-3 milli (m) and 10-6 micro (μ). The copyright© in this document belongs to the Crown Agents and this document should not be reproduced in whole or in part or used for any purpose without the express written permission of the Crown Agents.
1. | General Description: Defibrillator
| (1) | To export specification suitable for operation in Georgia
| (2) | Conformity of Medical Devices Approval
The equipment shall comply with the Pharmaceutical and Medical Devices Act of Japan. The Notification, Certification or Approval Number according to the Classification of Medical Devices shall be specified by bidder. A copy of the Certificate shall be submitted together with offered catalogue. If the number is shown on catalogue, no need to submit.
| (3) | Reference Model: TEC-5621K
2. | Purpose: To recover function of patient's heart by providing direct current.
3. | Composition (for one unit)
| (1) | Main unit (Defibrillator) |  |  |  | 1 unit
| (2) | Accessories
|  | 1) | Standard Accessories |  |  | 1 set
|  |  | ① | External paddle |  | 1 set
|  |  | ② | External paddle holder |  | 1 set
|  |  | ③ | ECG connection cord |  | 1 set
|  |  | ④ | Recording paper (roll, 50 mm x 30 m) |  | 1 pc.
|  | 2) | Rechargeable battery |  |  | 1 pc.
|  | 3) | SD card |  |  | 1 set
|  | 4) | Wall mount |  |  | 1 set
4. | Specifications
| (1) | Defibrillator
|  | 1) | Output energy |  |  | 2 to 200 J and 270 J
|  | 2) | Waveform |  |  | Biphasic effective defibrillation
|  | 3) | Charging time to 200 J |  |  | less than 5 secs.
|  | 4) | Manual mode and AED mode |  |  | to be provided
|  | 5) | Discharge mode |  |  | Synchronized
| (2) | ECG amplifier
|  | 1) | Electrode lead |  |  | 3-lead or more
|  | 2) | Heart rate |  |  | 25 to 300 bpm or wider
| (3) | Monitor
|  | 1) | Display parameters |  |  | Heart rate, SpO2 (%), Pulse rate, ETCO2, Respiration rate
|  | 2) | No. of trace |  |  | max. 4 traces
|  | 3) | Display type and size |  |  | Color LCD, 6.5" or more
|  | 4) | Sweep speed |  |  | 25 mm/sec. or more
|  | 5) | Alarm function |  |  | Audible and visual
| (4) | Printer |  |  |  | Built-in
|  | 1) | Paper speed |  |  | 25mm/sec. or more
| (5) | Battery
|  | 1) | Battery capacity |  |  | 12V, 2,800 mAh
|  | 2) | Operation time |  |  | Minimum 100 times discharges at 270 J
| (6) | Wall mount
|  | 1) | Wall mount device with fixing belt |  |  | to be provided
| (7) | Safety Standard |  |  |  | Class I, Type CF
| (8) | Mechanical strength: MIL-STD |  |  |  | to be specified
| (9) | Power requirement |  |  |  | AC220V ±10%, 50Hz, Single phase and Rechargeable
battery
| (10) | Type of AC plug |  |  |  | C type compatible
5. | Spare Parts (per unit)
| (1) | ECG electrode Lead, 3 electrodes |  |  |  | 1 pc.
| (2) | ECG electrode Lead, 6 electrodes |  |  |  | 1 pc.
| (3) | ECG connection cord |  |  |  | 1 pc.
| (4) | AED disposable pads (1 pair/set, for adult/pediatric, 1.5m) |  |  |  | 1 set
| (5) | AED dispo-pad adapter cable |  |  |  | 1 pc.
| (6) | Multi parameter/SpO2 unit |  |  |  | 1 set
| (7) | SpO2 Finger probe for adult |  |  |  | 1 pc.
| (8) | SpO2 multi site probe, for neonate child adult |  |  |  | 1 pc.
| (9) | SpO2 connection cord |  |  |  | 1 pc.
6. | Consumables (per unit)
| (1) | ECG disposable electrode adult (150pcs./set) |  |  |  | 1 set
| (2) | ECG disposable electrode child (150pcs./set) |  |  |  | 1 set
| (3) | AED disposable pads (1 pair/set, for adult/pediatric, 1.5m) |  |  |  | 2 sets
| (4) | Gel (200 g) |  |  |  | 1 pc.
| (5) | Recording paper (roll, 50 mm x 30 m) (10 pcs./set) |  |  |  | 1 set
7. | Conformity of Standards
The equipment shall be conformed to the Standards of CE, TUV, IEC, ISO9001, ISO13485 and /or JIS. (A copy of the Certificate shall be submitted together with offered catalogue. If conformity of the standard is shown on catalogues, no need to submit.)
8. | Model (s) details
| (1) | Defibrillator
|  | 1) Manufacturer |  |  |  | Please state for the model offered.
|  | 2) Model Number |  |  |  | Please state for the model offered.
|  | 3) Country of Origin |  |  |  | to be specified
9. | Logo display: Japanese national flag sticker
| (1) | To be affixed on each component of the main unit and accessories wherever possible.
10. | Warranty and Aftersales Service and Support
| (1) | Please provide details of the warranty conditions that will apply to this offer (Minimum 1 year).
| (2) | Full details of your local agent including:
|  | 1) | Name
|  | 2) | Address
|  | 3) | Phone
|  | 4) | Fax
|  | 5) | E-mail
|  | 6) | Web-site
| (3) | Please provide details how local agent will honour warranty claims, noting the local agent shall provide End-users with support including inspection, repair, maintenance, and/or supply of spare-parts (if applicable and as per manufacturer's recommendation) during the warranty period upon the request of the End Users.
| (4) | Duration of appointment of local agent : please specify appointed year and duration of appointment (e.x. appointed from 20XX, duration: no. of years)
| (5) | Numbers and types of technical staff employed.
| (6) | Please indicate the spares that will be held at the local agent’s premises for the equipment on offer, and the approximate value (as per Manufacturer's recommendation).
| (7) | Please confirm that the spare parts (if applicable and as per Manufacturer's recommendation) will be available for a period of minimum of 5 years from handover of Goods (7 years is desirable).
| (8) | Please state if the local agent can provide End-users with support including inspection, repair, maintenance, and/or supply of spare-parts (if applicable and as per Manufacturer's recommendation) after the warranty period upon the request of End-users outside of the scope of the resulting Contract of this Tender. Please provide the details of the support they can offer in this regard.
| (9) | Note that following delivery under any contract for supply, the bidder shall be responsible for ensuring that the Principal and End-User are provided with any service updates, service bulletins, safety notices, etc. either direct or via the local agent.
11. | O&M Manuals
Instruction manuals for operation and maintenance for the Goods. Manuals shall be English and Russian translation if possible. The end-user shall receive a set of manuals for each unit.
12. | Spare-parts list
One (1) set of Manufacturer’s List for recommended standard spare-parts or catalogue shall be supplied. The list is preferable to include as much as detailed information such as price, description, origin, serial number, etc. as is practically possible, but manufacture’s standard format is acceptable. In case it is not possible to provide such spare-parts list or catalogue, please state the reasons.

## Electrocardiograph
Lot No., Item No. | [6 empty] | Bidder's reference:
CA Reference |  |  |  |  | GEN20/109741D/0001 | Revision No.
CA Specification | [5 empty] | Revision Date
Prepared by |  |  |  |  | HK | Bidder's reference:
DETAILED SPECIFICATION | [6 empty] | STATEMENT of COMPLAIANCE (*)
Instructions to Bidder: Bidders are to offer a standard production model most closely matching the specification below and provide details of the offer. * Bidders must enter "Comply" or "Not comply" against each paragraph of this specification, comment as necessary, and sign and stamp each page. Where information is presented in the specification, the comment should be “Noted and Understood”. Failure to complete this statement of compliance may result in the offer being rejected. In the event of electronic bidding by E-mail or by means of scanning of the Specification for completion by word processing, the specification shall not be altered in any way. Any deviation to specification should be stated and if necessary, fully explained as a comment in the compliance column without making any changes to the specification. Any changes to the specification may invalidate the Bid. Each compliance sheet shall be duly authorised, signed, dated and stamped using the bidders company stamp. | [6 empty] | UNITS ARE TO BE STATED USING THE SI SYSTEM. For example length in metres (m). Multiples and sub-multiples to be used are: 106 mega (M); 103 kilo (k); 10-3 milli (m) and 10-6 micro (μ). The copyright© in this document belongs to the Crown Agents and this document should not be reproduced in whole or in part or used for any purpose without the express written permission of the Crown Agents.
1. | General Description: Electrocardiograph
| (1) | To export specification suitable for operation in Georgia
| (2) | Conformity of Medical Devices Approval
The equipment shall comply with the Pharmaceuticals and Medical Devices Act of Japan. The Notification, Certification or Approval Number according to the Classification of Medical Devices shall be specified by bidder. A copy of the Certificate shall be submitted together with offered catalogue. If the number is shown on catalogue, no need to submit.
| (3) | Reference Model: FX-8400
2. | Purpose: |  |  | for detecting cardiac problems in physical examinations and for monitoring the patient's condition during and after surgery or intensive care.
3. | Composition (for one unit)
| (1) | Main unit (ECG) |  |  |  | 1 unit
| (2) | Accessories
|  | 1) | Standard Accessories |  |  | 1 set
|  |  | ① | Patient Cable |  | 1 set
|  |  | ② | Chest Electrode for Adult (6 pcs./set) |  | 1 set
|  |  | ③ | Limb Electrode for Adult (4 pcs./set) |  | 1 set
|  | 2) | Rechargeable Battery |  |  | 1 pc.
|  | 3) | Recording Paper (roll width 210 mm x 30 m) |  |  | 1 pc.
|  | 4) | ECG gel |  |  | 1 pc.
|  | 5) | Carry case |  |  | 1 pc.
4. | Specifications
| (1) | Electrocardiograph
|  | 1) | Type |  |  | 3 ch., 6 ch., 12 ch.
|  | 2) | Sensitivity Selection |  |  | Auto, 1/4, 1/2, 1, 2
|  | 3) | Leads |  |  | Standard 12-lead
|  | 4) | Differential and Common-Mode Offset Voltage (Electrode-Skin Voltage) |  |  | ±600 mV and more
|  | 5) | Sine Wave Characteristics |  |  | 0.05 Hz to 250 Hz
|  | 6) | Low Frequency Characteristics (Time Constant) |  |  | 3.2 sec. and more
|  | 7) | Common Mode Signal Suppression |  |  | 103 dB and above (2 mm [p-v] and below at sensitivity level 1)
|  | 8) | Filter |  |  | AC filter : ₋ 20dB or less at 50 Hz or 60 Hz Muscle filter : ₋ 3dB ( ₋ 6 dB/oct) or less at 35 Hz or 25 Hz Drift : ₋3 dB (₋ 12 dB/oct) or less at 0.25 Hz or 0.5 Hz
|  | 9) | Printing Method |  |  | Thermal Print Head
|  | 10) | Printing Speed |  |  | 5, 10, 12.5, 25, 50 mm/s ±2% or less
|  | 11) | Printing Channel |  |  | 3 ch., 6 ch., 12 ch.
|  | 12) | Printing Paper |  |  | Roll paper with grids, 210mm Z-fold paper with grids, 210mm
|  | 13) | Display |  |  | 8” Color LCD, 800 × 480 dots (with LED backlight)
|  | 14) | A/D Conversion |  |  | 24-bit
|  | 15) | Sampling Rate |  |  | 8,000 samples/sec.
|  | 16) | LAN Port |  |  | Conforms to IEEE802.3u 100BASE-TX (The cable must be within 50 m.)
|  | 17) | USB Port |  |  | Compatible with USB2.0 Full Speed, 3 ports
|  | 18) | SD Card |  |  | Compatible with SD Card Specification 2.0
|  | 19) | Serial Port |  |  | Compatible with RS-232C, 2 ports
|  | 20) | Dimensions |  |  | Approximately 330 mm (W) × 350 mm (D) × 85 mm (H) (not including the protrusion)
|  | 21) | Safety Standard |  |  | Class I, Type CF
| (2) | Rechargeable Battery
|  | 1) | Battery |  |  | DC 14.8V
|  | 2) | Operation time |  |  | 240 min.
| (3) | Power requirement |  |  |  | AC220V ±10%, 50Hz, Single phase and Rechargeable battery
| (4) | Type of AC plug |  |  |  | C type compatible
5. | Spare Parts
Nil
6. | Consumables (per unit)
| (1) | Patient Cable |  |  |  | 1 set
| (2) | Chest Electrode for Adult (6 pcs./set) |  | Chest Electrode for Adult (6 pcs./set) |  | 1 set
| (3) | Limb Electrode for Adult (4 pcs./set) |  | Limb Electrode for Adult (4 pcs./set) |  | 1 set
| (4) | Recording Paper (roll width 210 mm x 30 m) |  |  |  | 1 pc.
| (5) | ECG gel |  |  |  | 1 pc.
7. | Conformity of Standards
The equipment shall be conformed to the Standards of CE, TUV, IEC, ISO9001, ISO13485 and /or JIS. (A copy of the Certificate shall be submitted together with offered catalogue. If conformity of the standard is shown on catalogues, no need to submit.)
8. | Model (s) details
| (1) | Electrocardiograph
|  | 1) Manufacturer |  |  |  | Please state for the model offered.
|  | 2) Model Number |  |  |  | Please state for the model offered.
|  | 3) Country of Origin |  |  |  | to be specified
9. | Logo display: Japanese national flag sticker
| (1) | To be affixed on each component of the main unit and accessories wherever possible.
10. | Warranty and Aftersales Service and Support
| (1) | Please provide details of the warranty conditions that will apply to this offer (Minimum 1 year).
| (2) | Full details of your local agent including:
|  | 1) | Name
|  | 2) | Address
|  | 3) | Phone
|  | 4) | Fax
|  | 5) | E-mail
|  | 6) | Web-site
| (3) | Please provide details how local agent will honour warranty claims, noting the local agent shall provide End-users with support including inspection, repair, maintenance, and/or supply of spare-parts (if applicable and as per manufacturer's recommendation) during the warranty period upon the request of the End Users.
| (4) | Duration of appointment of local agent: please specify appointed year and duration of appointment (e.x. appointed from 20XX, duration: no. of years)
| (5) | Numbers and types of technical staff employed.
| (6) | Please indicate the spares that will be held at the local agent’s premises for the equipment on offer, and the approximate value (as per Manufacturer's recommendation).
| (7) | Please confirm that the spare parts (if applicable and as per Manufacturer's recommendation) will be available for a period of minimum of 5 years from handover of Goods (7 years is desirable).
| (8) | Please state if the local agent can provide End-users with support including inspection, repair, maintenance, and/or supply of spare-parts (if applicable and as per Manufacturer's recommendation) after the warranty period upon the request of End-users outside of the scope of the resulting Contract of this Tender. Please provide the details of the support they can offer in this regard.
| (9) | Note that following delivery under any contract for supply, the bidder shall be responsible for ensuring that the Principal and End-User are provided with any service updates, service bulletins, safety notices, etc. either direct or via the local agent.
11. | O&M Manuals
Instruction manuals for operation and maintenance for the Goods. Manuals shall be English and Russian translation if possible. The end-user shall receive a set of manuals for each unit.
12. | Spare-parts list
One (1) set of Manufacturer’s List for recommended standard spare-parts or catalogue shall be supplied. The list is preferable to include as much as detailed information such as price, description, origin, serial number, etc. as is practically passible, but manufacture’s standard format is acceptable. In case it is not possible to provide such spare-parts list or catalogue, please state the reasons.